Acne treatment procedure

with the JETT PLASMA Medical II device

## PRE-TREATMENT PROCEDURE

- Exclusion criteria

Before the procedure, it is necessary to find out about any contraindications that would preclude treatment (peacemaker, Holter ecg monitoring system; another implanted electrical device; epilepsy; pregnancy; metal implants in the treated area; skin diseases and skin inflammations (except acne; physician evaluation); acute inflammatory disease (except acne); any untreated/badly treated disease in treatment area; oncological disease in the place of treatment; allergy to local anaesthetics; allergy to disinfectants)

- Acne examination

- We will assess the nature and size of the acne

- We check the condition of the skin

- Disinfection and application of local anesthesia

- Make-up is removed

- Proper skin disinfection is performed (given that the JETT PLASMA Medical II device works with electricity, we recommend an alcohol-free disinfectant)

## TREATMENT PROCEDURE

- Setting up the JETT PLASMA Medical II device

- For treatment, a flat applicator 3 or 5 mm is selected depending on the size of the treated layer.

- The device is turned on using the power switch on the back of the device.

- The patient is connected to the device. First, the patient is connected (using a cylinder or disposable electrode) to the device and then the cable to the device.

- After the patient is connected to the device, contact with the patient and contact with the applicator should be avoided.

- Patient preparation

- The patient is placed in a horizontal position.

- The position of the body part with acne and the lighting are checked, which should be as comfortable as possible for both the patient and the doctor.

- The skin is disinfected again with alcohol-free disinfectant.

- Before starting the treatment, wait approximately 1 minute for the disinfectant to evaporate.

- Treatment

- Sterile gloves are put on.

- Select the DERMATOLOGY-SILVER APPLICATORS mode on the device and set the intensity to 5 (the intensity can be reduced based on the patient's sensitivity)

- A layer of conductive indifferent gel or any serum suitable for acne-prone skin is applied to the treated area.

- The applicator is placed on the treated area and gently run over the skin.

- The recommended treatment time is around 15 – 20 minutes depending on the size of the treated area.

- After the treatment with the flat applicator, flat applicator is replaced with a golden applicator 14.4 mm or 24.4 mm.

- Set DERMATOLOGY - GOLDEN APPLICATORS mode on the device and set the intensity to 5.

- The discharge is applied in the form of a scanning, which means the discharge is quickly passed over the treated area, thereby heating the skin.

- The scanning result is gray skin

## POST-PROCEDURE TREATMENT

- After the procedure, clean the skin with a disinfectant; under no circumstances should you apply regular corrective cosmetics.

- Serum for acne-prone skin is applied to the treated area.

- Protect your skin from the sun by using SPF 50+ creams.

A minimum of 7 sessions is recommended with the following spacing:

After the first treatment, the 2nd treatment is done 3-4 days apart, the 3rd and 4th treatments are done weekly, the 5th and 6th treatments are done two weeks apart, and the remaining treatments are always done monthly.