Treatment procedure – melasma

with JETT PLASMA Medical II

It should be first evaluated, whether it is really pigmentation and not any inflammatory change of skin colour or another skin disorder.

Check, whether they are the malign or pre-malign conditions of lesions or not. It is important to focus on a differential diagnosis of basaloma, squamous cell carcinoma and melanoma.

## THE PRE-TREATMENT PROCEDURE

- Exclusion criteria for subject selection

Before the procedure, it is necessary to find out about any contraindications that would preclude treatment (peacemaker, Holter ecg monitoring system; another implanted electrical device; epilepsy; pregnancy; metal implants in the treated area; skin diseases and skin inflammations (except acne; physician evaluation); acute inflammatory disease (except acne); any untreated/badly treated disease in treatment area; oncological disease in the place of treatment; allergy to local anaesthetics; allergy to disinfectants)

- The examination of a lesion

- Evaluate the character of affection;

- Check the skin condition;

- Then, include the patients with pigmentations into the indicated conditions.

- The disinfection and the application of local anaesthesia

- Remove the make-up;

- Disinfect thoroughly the skin (with regard to the fact that JETT PLASMA Medical II works with electric energy, we recommend to use an alcohol-free disinfectant);

- Anaesthesia application:

- In case of small layers, apply anaesthesia in the form of the anaesthetic cream with a minimum content of 5% Lidocaine or Tetracaine (e.g.: Emla cream or Lidocaine gel);

- Wait for the cream effect for minimum 20 minutes;

- It is suitable to use occlusion to speed up the effect of anaesthesia (cover the applied cream with an airtight foil – plastic foil);

- In case of larger layers, apply the local anaesthesia in the form of an injection (Supracaine 4%, Mesocaine 1%, Lidocaine, etc.)

## THE TREATMENT PROCEDURE

- The setting of JETT PLASMA Medical II

- For treatment, we choose the golden applicator 14.4/24.4 mm, which we attach to the Plasma Pen

- The device is turned on using the power switch on the back of the device.

- The patient is connected to the device. First, the patient is connected (using a cylinder or disposable electrode) to the device and then the cable to the device

- After the patient is connected to the device, contact with the patient and contact with the applicator should be avoided.

- The preparation of a patient

- Check the position of the body part and the lighting which should be most comfortable both for the patient and the doctor;

- Disinfect again the skin using an alcohol-free disinfectant;

- Wait roughly 1 minute prior to start the treatment until the disinfection is evaporated.

- Treatment

- Put on a pair of sterile gloves;

- Select the DERMATOLOGY - GOLDEN APPLICATORS mode

- Set the intensity on the device to 4 or 5 (the choice of intensity depends on our subjective consideration, for better results we choose level 5, in the case of a sensitive patient, level 4)

- Get closer to the patient´s skin with the head tip and as soon as you get to 4 mm above the skin surface, the spark discharge will ignite;

- If the tip of the head remains in a correct height above the skin, the plasma discharge will be continuous;

- Remove the pigmentation by scanning gradually its whole area place by place;

- Always wipe away the carbonised layer with cellulose pads with physiological solution or disinfecting solution (free of spirit!!!).

- If the pigmentation persists, repeat the whole procedure - the point is that it cannot be presumed if the pigmentation is only epidermal or also dermal.

- After the treatment, disinfect thoroughly the whole treated area.

## AFTER-TREATMENT THERAPY

- After the treatment, clean the skin with a disinfectant and cover it by a plaster as the case may be. (Keep the plaster maximum until the evening). By no means may you apply common corrective cosmetics.

- The skin can be washed up in the evening after the removal;

- Disinfect the treated places two times a day for 2 – 3 days after the treatment;

- After disinfection, it is suitable to apply a special regeneration gel containing hyaluronic acid above 1 % intended for an injured skin twice a day for the next 5 – 7 days;

- After the regeneration gel dries up (approx. 10 min.), corrective cosmetics (make-up) can be applied on the skin;

- Protect the skin from solar radiation by applying the creams with SPF 50+ and sun glasses and head cover for a minimum of 7-10 days after the treatment;

If the treatment is done correctly as instructed, the hyperpigmentation should be removed immediately. Alternatively, a carbonised layer may remain on the skin - a crust - which will peel off within a few days.

Healing usually takes a week, rarely longer. If the hyperpigmentation removed is of a larger size, healing may be extended up to 14 days.