Treatment procedure –

The removal of atrophic, hypertrophic and keloid scars with JETT PLASMA Medical II

The type of scars should always be evaluated.

We distinguish the following types of scars:

- Atrophic, sunken under the skin level, e.g. scars after acne;

- Hypertrophic, raised above the surrounding epidermis. They are harder than the surrounding tissue, however they do not extend beyond the boundary of trauma;

- Keloid, overgrowing markedly the surrounding area, hard, shiny and whitish, extending beyond the boundary of trauma.

Never treat fresh scars. They can be treated first after two months from the injury.

Recommend your client to apply a silicon sheet or ointment to fresh scars at night on a long-term basis.

THE PRE-TREATMENT PROCEDURE

- Exclusion criteria for subject selection

Before the procedure, it is necessary to find out about any contraindications that would preclude treatment (peacemaker, Holter ecg monitoring system; another implanted electrical device; epilepsy; pregnancy; metal implants in the treated area; skin diseases and skin inflammations (except acne; physician evaluation); acute inflammatory disease (except acne); any untreated/badly treated disease in treatment area; oncological disease in the place of treatment; allergy to local anaesthetics; allergy to disinfectants)

- The examination of a scars

- Evaluate the scar character;

- Check the skin condition.

- The disinfection and the application of local anaesthesia

- Remove the make-up;

- Disinfect thoroughly the skin (with regard to the fact that JETT PLASMA Medical II works with electric energy, we recommend to use an alcohol-free disinfectant);

- Anaesthesia application:

    - In case of small scars, apply anaesthesia in the form of the anaesthetic cream with a minimum content of 5% Lidocaine or Tetracaine (e.g.: Emla cream or Lidocaine gel);

    - Wait for the cream effect for minimum 20 minutes;

    - It is suitable to use occlusion to speed up the effect of anaesthesia (cover the applied cream with an airtight foil – plastic foil);

    - In case of larger scars, apply the local anaesthesia in the form of an injection (Supracaine 4%, Mesocaine 1%, Lidocaine, etc.)

## THE TREATMENT PROCEDURE

- The setting of JETT PLASMA Medical II

- For treatment, choose the golden applicator 14.4/24.4 mm, which attach to the Plasma Pen

- The device is turned on using the power switch on the back of the device.

- The patient is connected to the device. First, the patient is connected (using a cylinder or disposable electrode) to the device and then the cable to the device

- After the patient is connected to the device, contact with the patient and contact with the applicator should be avoided.

- The preparation of a patient

- Place the patient in a horizontal position

- Check the position of the body part with the scar and the lighting, which should be as comfortable as possible for the patient and the doctor

- Disinfect the skin again with an alcohol-free disinfectant

- Wait approximately 1 minute for the disinfectant to evaporate before starting the treatment

- Treatment

- Put on a pair of sterile gloves;

- Select the DERMATOLOGY - GOLDEN APPLICATORS mode

- Set the intensity on the device to 4 or 5 (the choice of intensity depends on our subjective consideration, for better results we choose level 5, in the case of a sensitive patient, level 4)

- Get closer to the patient´s skin with the head tip and as soon as you get to 4 m above the skin surface, the spark discharge will ignite;

- If the tip of the head remains in a correct height above the skin, the plasma discharge will be continuous.

ATROPHIC SCARS

    - Treat the atrophic scar by scanning the whole area of the scar so as to have the area grey (not brown!!!).

HYPERTROPHIC AND KELOID SCARS

Note: In case of keloid scars, we recommend to treat first only a small area of approximately 2 x 2 mm as a trial and to evaluate the application effect after 3 – 4 weeks. Some keloid scars react to minor traumata (which is even the treatment of a small area), others react by further growing of ligament and hypertrophy and those are not suitable for Plasma treatment.

    - Treat the hypertrophic and keloid scars by layers so as to have the surface brown to black.

    - Always wipe away the carbonised layer with cellulose pads with physiological solution or disinfecting solution (free of alcohol!!!).

- You should always check what layer you have removed – it is necessary to treat the whole scar and to avoid getting under the level of the surrounding skin and causing thus any defect in the skin;

- We recommend to burn out the scar to only 1 – 2 mm above the level of the surrounding skin during the first treatment;

- You should keep in mind that the plasma application causes edema in the place of application due to high temperature;

- Further treatment of the scar can be performed after 14 days and the scar can be flattened to the level of the surrounding skin

- After the treatment is completed, carefully disinfect the entire treated area

## AFTER-TREATMENT THERAPY

- After the treatment, clean the skin with a disinfectant and cover it by a plaster as the case may be. (Keep the plaster maximum until the evening). By no means may you apply common corrective cosmetics.

- The skin can be washed up in the evening after the removal;

- Disinfect the treated places two times a day for 2 – 3 days after the treatment;

- After disinfection, it is suitable to apply a special regeneration gel containing hyaluronic acid above 1 % intended for an injured skin twice a day for more 5 – 7 days;

- After the week, grease the skin thoroughly with creams to prevent the crust from preliminary pulling-off;

- Protect the skin from solar radiation by applying the creams with SPF 50+ and sun glasses and head cover for a minimum of 7-10 days after the treatment;

- We recommend seeing the doctor after a week.

If the treatment was performed correctly according to the instructions, the scar should be removed after two sessions at the latest.

During the first treatment, the scar is burned only to the level of approximately 1 – 2 mm above the level of the surrounding skin. After 14 days, the second treatment will level the scar to the surrounding skin.

Healing takes usually one week. In extraordinary cases, it may take a longer time. If larger scar was removed, its healing may last up to 14 days.