Treatment procedure – Trichiasis, distichiasis

with the JETT PLASMA Medical II device

PRE-TREATMENT PROCEDURE

- Exclusion criteria

Before the procedure, it is necessary to find out about any contraindications that would preclude treatment (peacemaker, Holter ecg monitoring system; another implanted electrical device; epilepsy; pregnancy; metal implants in the treated area; skin diseases and skin inflammations (except acne; physician evaluation); acute inflammatory disease (except acne); any untreated/badly treated disease in treatment area; oncological disease in the place of treatment; allergy to local anaesthetics; allergy to disinfectants)

- Disinfection and application of local anesthesia

  - Standard preoperative preparation will take place in the operating room. After removing make-up, Benoxi gtt will be applied to both eyes,The skin is properly disinfected (given that the JETT PLASMA Medical II device works with electricity, an alcohol-free disinfectant is recommended) and covered with surgical drapes.

  - As part of preoperative preparation to increase comfort, it is advisable to apply a lidocaine cream to the patient's eyelids. Wait at least 20 minutes for the cream to take effect.

  - Local anesthesia is applied to the skin according to the workplace's practices (Supracaine 4%, Mesocaine 1%, Lidocaine, Mepivastesin, Ubivastesin, etc.). It is always applied with a 1ml insulin syringe.

TREATMENT PROCEDURE

# Setting up the JETT PLASMA Medical II device

- For the treatment, the ophthalmological Gold/Steel applicator is selected, or the 14.4/24.4 mm gold applicator, which is attached to the Plasma Pen.

- Plug the device into the socket and turn it on using the switch on the back of the device.

- Connect the appropriate connecting cables to the disposable electrode, which is then stuck to the patient's skin (the inner side of the forearm, where there is no hair, is suitable, the patient can just roll up the sleeve, or the back of the arm) or to the cylinder electrode, which is placed in the patient's hand and held firmly throughout the treatment.

- According to safety regulations, the connecting cable is only connected to the device now. This prevents contact of the applied applicator with the patient's body and with unprotected parts of the body.

- After the patient is connected to the device, avoid contact with the patient and the patient with the applicator.

# Patient preparation

- The head position and lighting are checked, which should be as comfortable as possible for both the patient and the doctor.

- The skin is disinfected again with alcohol-free disinfectant.

- Before starting the treatment, wait approximately 1 minute for the disinfectant to evaporate.

- It is also important to cover the bulb itself with a non-metallic corneal cap. The smoke from the evaporation and the discharge itself are unpleasant, and without covering the bulb, a painful burn of the cornea could occur.

# Treatment

- Sterile gloves are put on.

- The device is set to level 6 to avoid burning the eyelid margin too much. The choice of level depends on the doctor's subjective judgment; for better results, it can be gradually set to level 8. The eyelid is everted and, using a slit lamp, the tip of the ophthalmological applicator is inserted as deeply as possible into the hair canal, and then the device is turned on for 1–2 seconds..This will destroy the hair bulb.

- Disconnects the appropriate connection cable from the device, the disposable electrode is peeled off from the patient's body or the cylinder electrode is handed to the treating worker.

POST-PROCEDURE TREATMENT

- After the procedure, antibiotic ointment with glucocorticoids Maxitrol ung is applied.

- It is important to protect the eyes and their surrounding area from sunlight by using SPF 50+ creams and wearing sunglasses and a hat for at least 10 days after the procedure.

Possible skin deformation caused by local anesthesia and edema from skin heating at the application site must always be taken into account.